## mya-archiver
HPS SVT MYA Variables
Notes:
"stat" variables are ON/OFF booleans
|  | Deadband |  | Alarm
|  |  |  | severe | warning | warning | severe
Category | Variable | low | high | low | low | high | high
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:avdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:avdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:avdd:i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.59 | 0.61
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:avdd:i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:avdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:avdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:dvdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:dvdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:dvdd:i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:dvdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:dvdd:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:v125:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:v125:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:v125:i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:v125:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:v125:i_rd | 0.01 | 0.01 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:dvdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:dvdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:dvdd:v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:dvdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:dvdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:avdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:avdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:avdd:v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:avdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:avdd:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:v125:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:v125:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:v125:v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:v125:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:v125:v_set_rd | 1 | 1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:dvdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:dvdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:dvdd:vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:dvdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:dvdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:avdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:avdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:avdd:vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:avdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:avdd:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:v125:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:v125:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:v125:vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:v125:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:v125:vn | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:avdd:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:avdd:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:avdd:vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:avdd:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:avdd:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:v125:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:v125:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:v125:vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:v125:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:v125:vf | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:dvdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:dvdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:dvdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:dvdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:dvdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:avdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:avdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:avdd:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:avdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:avdd:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/
/SVT/Voltages/Low voltages/ | SVT:lv:0:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:0:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:1:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:2:2:v125:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:2:3:v125:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:3:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:3:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:4:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:5:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:6:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:7:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:2:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:8:3:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:0:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:1:v125:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:9:2:v125:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Voltages/Low voltages/ | SVT:lv:9:3:v125:stat | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Temperatures/ | SVT:temp:hyb:0:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:0:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:0:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:0:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:1:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:1:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:1:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:1:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:2:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:2:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:2:2:temp0:t_rd | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Temperatures/ | SVT:temp:hyb:2:3:temp0:t_rd | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Temperatures/ | SVT:temp:hyb:3:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:3:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:3:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:3:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:4:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:4:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:4:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:4:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:5:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:5:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:5:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:5:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:6:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:6:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:6:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:6:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:7:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:7:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:7:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:7:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:8:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:8:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:8:2:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:8:3:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:9:0:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:9:1:temp0:t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
/SVT/Temperatures/ | SVT:temp:hyb:9:2:temp0:t_rd | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Temperatures/ | SVT:temp:hyb:9:3:temp0:t_rd | 0.1 | 0.1 | - | - | - | - | No hybrid
/SVT/Temperatures/ | SVT:temp:fe:0:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:1:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:2:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:3:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:4:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:5:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:6:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:7:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:8:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:9:axixadc:t_rd | 1 | 1 | 35 | 37 | 45 | 47
/SVT/Temperatures/ | SVT:temp:fe:0:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:1:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:2:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:3:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:4:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:5:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:6:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:7:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:8:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:9:FebTemp0:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:0:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:1:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:2:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:3:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:4:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:5:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:6:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:7:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:8:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Temperatures/ | SVT:temp:fe:9:FebTemp1:t_rd | 1 | 1 | 20 | 22 | 28 | 30
/SVT/Daq | SVT:lv:*:*:sync:sync_rd | 1 | 1 | 30 | 30 | 32 | 32
/SVT/Daq | SVT:daq:*:*:*:syncbase_rd | 10 | 10 | - | - | - | -
/SVT/Daq | SVT:daq:*:*:*:syncpeak_rd | 10 | 10 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:febnum | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:hybnum | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:rxcellerrorcount | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:rxframecount | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:rxframeerrorcount | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:rxlinkdowncount | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:rxlinkerrorcount | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:*:*:rxphyready | 1 | 1 | - | - | - | -
/SVT/Daq | SVT:daq:dpm:0:eventstate | 1 | 1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:digi:i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:anap:i_rd | 0.01 | 0.01 | 1.95 | 2.02 | 2.15 | 2.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:anap:i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:anap:i_rd | 0.01 | 0.01 | 1.95 | 2.02 | 2.15 | 2.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:anan:i_rd | 0.01 | 0.01 | 0.27 | 0.275 | 0.295 | 0.305
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:anan:i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:anan:i_rd | 0.01 | 0.01 | 0.27 | 0.275 | 0.295 | 0.305
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:digi:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:digi:v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:anap:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:anap:v_term | 0.01 | 0.01 | 6.5 | 6.55 | 6.75 | 6.8
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:anap:v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:anap:v_term | 0.01 | 0.01 | 6.5 | 6.55 | 6.75 | 6.8
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:anan:v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:0:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:1:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:2:anan:v_term | 0.01 | 0.01 | 6.25 | 6.3 | 6.55 | 6.6
/SVT/Voltages/Low voltages/ | SVT:lv:fe:3:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:4:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:5:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:6:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:7:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:8:anan:v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
/SVT/Voltages/Low voltages/ | SVT:lv:fe:9:anan:v_term | 0.01 | 0.01 | 6.25 | 6.3 | 6.55 | 6.6
/SVT/Voltages/High voltages/ | SVT:bias:bot:20:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:21:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:22:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:23:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:24:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:25:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:26:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:27:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:28:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:29:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:30:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:31:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:32:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:33:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:34:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:35:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:36:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:37:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:0:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:10:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:11:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:12:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:13:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:14:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:15:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:16:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:17:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:1:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:2:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:3:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:4:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:5:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:6:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:7:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:8:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:9:v_term | 0.01 | 0.01 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:20:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:21:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:22:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:23:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:24:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:25:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:26:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:27:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:28:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:29:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:30:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:31:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:32:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:33:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:34:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:35:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:36:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:37:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:0:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:10:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:11:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:12:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:13:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:14:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:15:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:16:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:17:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:1:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:2:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:3:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:4:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:5:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:6:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:7:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:8:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:9:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:20:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:21:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:22:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:23:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:24:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:25:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:26:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:27:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:28:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:29:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:30:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:31:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:32:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:33:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:34:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:35:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:36:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:bot:37:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:0:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:10:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:11:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:12:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:13:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:14:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:15:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:16:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:17:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:1:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:2:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:3:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:4:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:5:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:6:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:7:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:8:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/High voltages/ | SVT:bias:top:9:i_rd | 0.001 | 0.001 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:fl:0:i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.7 | 0.75
/SVT/Voltages/Low voltages/ | SVT:lv:fl:1:i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.7 | 0.75
/SVT/Voltages/Low voltages/ | SVT:lv:fl:2:i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.7 | 0.75
/SVT/Voltages/Low voltages/ | SVT:lv:fl:3:i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.7 | 0.75
/SVT/Voltages/Low voltages/ | SVT:lv:fl:0:v_term | 0.1 | 0.1 | 6.2 | 6.3 | 6.9 | 7
/SVT/Voltages/Low voltages/ | SVT:lv:fl:1:v_term | 0.1 | 0.1 | 6.2 | 6.3 | 6.9 | 7
/SVT/Voltages/Low voltages/ | SVT:lv:fl:2:v_term | 0.1 | 0.1 | 6.2 | 6.3 | 6.9 | 7
/SVT/Voltages/Low voltages/ | SVT:lv:fl:3:v_term | 0.1 | 0.1 | 6.2 | 6.3 | 6.9 | 7
/SVT/Voltages/Low voltages/ | SVT:lv:fl:0:v_sens | 0.1 | 0.1 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fl:1:v_sens | 0.1 | 0.1 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fl:2:v_sens | 0.1 | 0.1 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fl:3:v_sens | 0.1 | 0.1 | 5.4 | 5.45 | 5.55 | 5.6
/SVT/Voltages/Low voltages/ | SVT:lv:fl:0:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:fl:1:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:fl:2:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Voltages/Low voltages/ | SVT:lv:fl:3:stat | 0.1 | 0.1 | - | - | - | -
/SVT/Interlock | HPS_SVT:PLC:i:Flow_FE-Value | 0.1 | 0.1 | - | - | - | -
/SVT/Interlock | HPS_SVT:PLC:i:Flow_SVT-Value | 0.1 | 0.1 | - | - | - | -
/SVT/Interlock | HPS_SVT:PLC:i:Vacuum-Value | 0.1 | 0.1 | - | - | - | -
/SVT/Temperatures | HPS_SVT:PLC:i:RTD_SVT_Supply-Value | 0.1 | 0.1 | -21 | -20 | -18.2 | -18
/SVT/Temperatures | HPS_SVT:PLC:i:RTD_SVT_Return-Value | 0.1 | 0.1 | -21 | -20 | -18 | -17.8
/SVT/Temperatures | HPS_SVT:PLC:i:RTD_FE_Supply-Value | 0.1 | 0.1 | 19.5 | 20 | 21 | 21.5
/SVT/Temperatures | HPS_SVT:PLC:i:RTD_FE_Return-Value | 0.1 | 0.1 | 21.5 | 22 | 23 | 23.5
/SVT/Temperatures | HPS_SVT:CHILLER:TEMP (set deadband at 0.5) | 0.5 | 0.1 | -24 | -22 | -18 | -16
/SVT/Temperatures | HPS_SVT:CHILLER:TEMP:RD_SP (no alarms) | 0.5 | 0.5 | - | - | - | -
/SVT/Temperatures | HPS_SVT:CHILLER:START_STOP (0 off, 1 on) | 0.1 | 0.5 | - | - | - | -
/SVT/Temperatures | HPS_SVT:CHILLER:PRES (alarm points 0.5, 1.0, 2.25, 2.5) | 0.1 | 0.1 | 0.5 | 1 | 2.25 | 2.5
/SVT/Temperatures | HPS_SVT:CHILLER:PUMP_STAGE (this is an integer; no alarm) | 0.1 | 0.1 | - | - | - | -
/SVT/Temperatures | HPS_FE:CHILLER:TEMP (set deadband at 0.5) | 0.5 | 0.1 | 16 | 18 | 22 | 24
/SVT/Temperatures | HPS_FE:CHILLER:TEMP:RD_SP (no alarms) | 0.1 | 0.5 | - | - | - | -

## limits
|  | Deadband |  | Alarm
|  |  |  | severe | warning | warning | severe
Channel type | Channel | low | high | low | low | high | high
Hybrid DVDD | i_rd | 0.01 | 0.01 | 0.24 | 0.26 | 0.295 | 0.31
Hybrid AVDD | i_rd | 0.01 | 0.01 | 0.345 | 0.355 | 0.43 | 0.45
Also note that AVDD Feb 6, Hybrid 2 is much higher than normal. I'm not sure what's causing this, but it works fine regardless of the high current. That said, the alarm limits should be set higher than the rest. This is variable SVT:lv:6:2avdd:i_rd. | i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.59 | 0.61
Hybrid V125 | i_rd | 0.01 | 0.01 | 0.295 | 0.305 | 0.34 | 0.36
Hybrid DVDD | v_set_rd | 1 | 1 | 140 | 145 | 155 | 160
Hybrid AVDD | v_set_rd | 1 | 1 | 160 | 165 | 175 | 180
Hybrid V125 | v_set_rd | 1 | 1 | 170 | 175 | 185 | 190
Hybrid DVDD | vn | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
Hybrid AVDD | vn | 0.1 | 0.1 | 2.5 | 2.6 | 2.8 | 2.9
Hybrid V125 | vn | 0.1 | 0.1 | 1.15 | 1.25 | 1.45 | 1.55
Hybrid AVDD | vf | 0.1 | 0.1 | 2.4 | 2.5 | 2.7 | 2.8
Hybrid V125 | vf | 0.1 | 0.1 | 1.07 | 1.17 | 1.37 | 1.47
Hybrid DVDD | stat | 0.1 | 0.1 | - | - | - | -
Hybrid AVDD | stat | 0.1 | 0.1 | - | - | - | -
Hybrid V125 | stat | 0.1 | 0.1 | - | - | - | -
Hybrid bias | v_term | 0.01 | 0.01 | - | - | - | -
Hybrid bias | stat | 0.1 | 0.1 | - | - | - | -
Hybrid bias | i_rd | 0.001 | 0.001 | - | - | - | - | Dynamic scale not good enough to measure uA.
FEB Axi temperature | t_rd | 1 | 1 | 35 | 37 | 45 | 47
FEB temperature | t_rd | 1 | 1 | 20 | 22 | 28 | 30
Hybrid temperature | t_rd | 0.1 | 0.1 | -14 | -14.5 | -16 | -16.5
Hybrid sync | sync_rd | 1 | 1 | 30 | 30 | 32 | 32
Hybrid sync | syncbase_rd | 50 | 50 | - | - | - | -
Hybrid sync | syncpeak_rd | 50 | 50 | - | - | - | -
| SVT:daq:dpm:0:0:febnum | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:hybnum | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:rxcellerrorcount | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:rxframecount | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:rxframeerrorcount | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:rxlinkdowncount | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:rxlinkerrorcount | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:0:rxphyready | 1 | 1 | - | - | - | -
| SVT:daq:dpm:0:eventstate | 1 | 1 | - | - | - | -
FEB DIG | i_rd | 0.01 | 0.01 | 0.9 | 1 | 1.2 | 1.3
FEB ANAP 4HYB | i_rd | 0.01 | 0.01 | 2.9 | 3 | 3.2 | 3.3
FEB ANAP 2HYB | i_rd | 0.01 | 0.01 | 1.95 | 2.02 | 2.15 | 2.2
FEB ANAN 4HYB | i_rd | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34 |  | 0.01 | 0.01 | 0.3 | 0.31 | 0.33 | 0.34
FEB ANAN 2HYB | i_rd | 0.01 | 0.01 | 0.27 | 0.275 | 0.295 | 0.305
FEB DIG | v_term | 0.01 | 0.01 | 6.8 | 6.9 | 7.1 | 7.2
FEB ANAP 4HYB | v_term | 0.01 | 0.01 | 7.15 | 7.21 | 7.29 | 7.35
FEB ANAP 2HYB | v_term | 0.01 | 0.01 | 6.5 | 6.55 | 6.75 | 6.8
FEB ANAN 4HYB | v_term | 0.01 | 0.01 | 6.45 | 6.51 | 6.7 | 6.75
FEB ANAN 2HYB | v_term | 0.01 | 0.01 | 6.25 | 6.3 | 6.55 | 6.6
FEB DIG | v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
FEB ANAP | v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
FEB ANAN | v_sense | 0.01 | 0.01 | 5.4 | 5.45 | 5.55 | 5.6
FEB DIG | stat | 0.1 | 0.1 | - | - | - | -
FEB ANAP | stat | 0.1 | 0.1 | - | - | - | -
FEB ANAN | stat | 0.1 | 0.1 | - | - | - | -
FL DIG | i_rd | 0.01 | 0.01 | 0.45 | 0.5 | 0.7 | 0.75
FL DIG | v_term | 0.1 | 0.1 | 6.2 | 6.3 | 6.9 | 7
FL DIG | v_sense | 0.1 | 0.1 | 5.4 | 5.45 | 5.55 | 5.6
FL DIG | stat | 0.1 | 0.1 | - | - | - | -
FEB flow switch | stat | 0.1 | 0.1 | - | - | - | -
SVT flow switch | stat | 0.1 | 0.1 | - | - | - | -
Vacuum | stat | 0.1 | 0.1 | - | - | - | -
HPS_SVT:PLC:i:RTD_SVT_Supply-Value |  | 0.1 | 0.1 | -21 | -20 | -18.2 | -18
HPS_SVT:PLC:i:RTD_SVT_Return-Value |  | 0.1 | 0.1 | -21 | -20 | -18 | -17.8
HPS_SVT:PLC:i:RTD_FE_Supply-Value |  | 0.1 | 0.1 | 19.5 | 20 | 21 | 21.5
HPS_SVT:PLC:i:RTD_FE_Return-Value |  | 0.1 | 0.1 | 21.5 | 22 | 23 | 23.5
HPS_SVT:CHILLER:TEMP (set deadband at 0.5) |  | 0.5 | 0.5 | -24 | -22 | -18 | -16
HPS_SVT:CHILLER:TEMP:RD_SP (no alarms) |  | 0.5 | 0.5 | - | - | - | -
HPS_SVT:CHILLER:START_STOP (0 off, 1 on) |  | 0.1 | 0.1 | - | - | - | -
HPS_SVT:CHILLER:PRES (alarm points 0.5, 1.0, 2.25, 2.5) |  | 0.1 | 0.1 | 0.5 | 1 | 2.25 | 2.5
HPS_SVT:CHILLER:PUMP_STAGE (this is an integer; no alarm) |  | 0.1 | 0.1 | - | - | - | -
HPS_FE:CHILLER:TEMP (set deadband at 0.5) |  | 0.5 | 0.5 | 16 | 18 | 22 | 24
HPS_FE:CHILLER:TEMP:RD_SP (no alarms) |  | 0.1 | 0.1 | - | - | - | -